High-Voltage Trigger Generator Assembly Guide

For the image-activated droplet sorting system

High-voltage warning: This module can generate kilovolt-level pulses. Build, test, and operate it only with proper insulation, enclosure, current limiting, grounding, discharge procedures, and trained supervision.

1. Purpose And System Role

The high-voltage trigger generator converts a low-voltage pulse into a short high-voltage pulse applied to the microfluidic chip electrodes. The pulse creates dielectrophoretic force that deflects target droplets into the sorted outlet.

Signal chain: Python sorting command -> Arduino digital output -> trigger driver -> DC-DC high-voltage boost converter -> chip electrodes.

2. Target Electrical Behavior

| Parameter | Target / Starting Value | Notes |
| --- | --- | --- |
| Arduino serial rate | 9600 baud | Must match the Python runtime and firmware. |
| Arduino output pulse | 4 ms | Default firmware pulse width; verify on pins 8/9 before connecting the driver. |
| Low-voltage trigger | TTL-level digital output | Use only the Arduino side on breadboards; keep high-voltage wiring separate. |
| High-voltage pulse | about 1.2 kV | Starting actuation voltage for the demonstrated chip and droplets. |
| Pulse repetition target | up to 100 Hz workflow | The high-voltage output must decay below the actuation threshold before the next event. |
| Observed timing offset | about 120 ms | Setup-dependent delay between image trigger and droplet deflection; recalibrate for every build. |

3. Minimal Bill Of Materials

| Subsystem | Item | Function | Selection Notes | Qty |
| --- | --- | --- | --- | --- |
| Control | Arduino-compatible MCU | Receives USB serial commands and generates timed digital pulses. | Use a board with stable USB serial and 5 V or 3.3 V logic compatible with the trigger driver. | 1 |
| Host interface | USB cable | Connects the workstation to Arduino. | Use a shielded, mechanically stable cable to avoid intermittent serial disconnects. | 1 |
| Low-voltage power | DC supply | Powers the trigger driver and/or boost converter input. | The demonstrated design uses a 12 V pulse into the boost stage. Choose current rating based on converter demand. | 1 |
| Triggering | Signal-triggered driver module | Switches the low-voltage input to the DC-DC boost converter when Arduino pulses. | Must accept Arduino logic-level input and switch quickly and repeatably. | 1 per output path |
| High-voltage generation | DC-DC boost converter | Generates the kV-level pulse for dielectrophoretic actuation. | Target about 1.2 kV output; select a module with suitable isolation, insulation, discharge behavior, and current limiting. | 1 |
| Low-voltage prototyping | Breadboard or custom PCB | Holds Arduino-side resistors and low-voltage wiring. | Do not route kV-level signals through a solderless breadboard. | 1 |
| Signal conditioning | Resistors / pull-downs | Stabilize trigger inputs and prevent floating states. | Select values for the driver input stage; verify idle-low behavior on an oscilloscope. | as needed |
| Interconnects | Jumper wires | Low-voltage connections between Arduino, driver, and supply. | Use short, strain-relieved low-voltage wiring. Keep separate from high-voltage wiring. | as needed |
| High-voltage wiring | Rated HV cable/connectors | Connect boost output to chip electrodes. | Use insulation rating above the maximum pulse voltage and prevent exposed conductors. | as needed |
| Safety enclosure | Insulated box / cover | Prevents accidental contact with energized conductors. | Include cable strain relief and labels for high-voltage terminals. | 1 |
| Measurement | Oscilloscope + HV probe | Verifies pulse timing, amplitude, and decay. | Use a probe rated for the maximum voltage and bandwidth of the pulse. | 1 set |
| Discharge/safety | Bleeder/discharge path and power disconnect | Provides safe shutdown and reduces residual charge. | Design according to local electrical safety rules and measured converter behavior. | 1 set |

4. Arduino Output Mapping

The two-class firmware uses pin 8 for class-1 droplets and pin 9 for class-2 droplets. Commands are one ASCII character sent by the Python virtual-gate logic.

| Command | Droplet Class | Output Pin | Default Delay Bin |
| --- | --- | --- | --- |
| 0 | Class 1 | 8 | near-gate bin, 1.0 ms delay |
| 1 | Class 1 | 8 | near-gate bin, 1.0 ms delay |
| 2 | Class 1 | 8 | late bin, 1.5 ms delay |
| 3 | Class 1 | 8 | late bin, 1.5 ms delay |
| 4 | Class 2 | 9 | near-gate bin, 1.0 ms delay |
| 5 | Class 2 | 9 | near-gate bin, 1.0 ms delay |
| 6 | Class 2 | 9 | late bin, 1.5 ms delay |
| 7 | Class 2 | 9 | late bin, 1.5 ms delay |

5. Assembly Notes

Keep all Arduino, USB, and trigger-driver logic wiring physically separated from high-voltage wiring.

Use the breadboard only for low-voltage prototyping. Move repeated-use builds to a secured PCB or terminal block.

Mount the boost converter and high-voltage output in an insulated enclosure before operation.

Label low-voltage input, Arduino trigger input, high-voltage output, ground return, and discharge points.

Add strain relief to wires connected to the microscope stage or chip holder.

Confirm that the high-voltage ground/reference is connected exactly as intended for the electrode design.

6. Bench Verification Before Connecting The Chip

Upload the Arduino sketch and disconnect the high-voltage stage.

Send test commands from the workstation and measure Arduino output pins with an oscilloscope.

Confirm each output pulse is 4 ms and returns cleanly to LOW.

Connect the trigger driver with the high-voltage output still disconnected from the chip.

Connect the high-voltage output to a rated dummy load and measure with a high-voltage probe.

Confirm pulse amplitude, pulse width, voltage decay, and absence of unintended retriggering.

Power down and discharge the module before changing wiring.

7. Integration With The Sorting Chip

Connect the positive high-voltage output to the positive liquid-metal electrode and connect the ground electrodes to the generator reference/ground as required by the electrode layout. Verify this wiring with power off before enabling the high-voltage module.

Start sorting at low droplet density. Increase flow speed only after target droplets consistently enter the sorted outlet and neighboring droplets are not disturbed.

8. Acceptance Checklist

Arduino serial communication is stable at 9600 baud.

Output pins pulse only when the Python sorter sends target commands.

Trigger-driver input is not floating when idle.

High-voltage output reaches the required pulse amplitude and decays before the next event.

No exposed high-voltage conductor is accessible during operation.

The chip sorts test droplets before biological samples are introduced.

Voltage, pulse width, ROI, virtual gate, flow rates, and firmware delay values are recorded for each run.

9. Maintenance And Version Control

Record module part numbers, converter settings, and measured output waveforms for each build.

Recheck pulse behavior after replacing any driver, converter, Arduino board, or power supply.

Revalidate sorting after changing chip geometry, tubing, microscope magnification, or droplet speed.

Keep a photo of the assembled wiring with labels for future troubleshooting.

Appendix: Reference Assembly Image

Reference image of the low-cost high-voltage trigger module and representative trigger/output waveform.